## BOQ
BOQ – BILL OF QUANTITIES AND BUDGET
Yoseftal Medical Center, Eilat – Inpatient ward · Nurse call (REV04)
Client: ACE BG · Advantecnia Date: 23/10/2026
| Currency | ₪ ILS |  | Yellow cells = fill in (unit prices). Blue text = inputs; black = formulas; green = link to another sheet.
| Contingency | 5%
| VAT (Israel) | 18%
# | Code | Description | Unit | Qty | Unit price | Total | Remark
ADVANTECNIA / ACE BG – NURSE CALL
1 | ACE82012153 | Red bed call panel + bayonet socket | pcs | 12 |  | 0 | 1 per bed · h 1.15 m
2 | ACE39060521 | Call pear handset (2 m cable to bayonet) | pcs | 12 |  | 0 | 1 per bed
3 | ACE82082137 | Presence / cancel | pcs | 6 |  | 0 | 4 rooms + 2 WCs · h 1.35 m
4 | EOP004 | Corridor light (red · green · blue) | pcs | 7 |  | 0 | 4 rooms + B3 ×2 + B2 · h 2.42 m
5 | ACE40320303 | SOS bathroom / shower pull-cord | pcs | 3 |  | 0 | B2 ×1 · B3 ×2 · h 1.85 m
6 | ECV169 | 24-pin PCB 12-48 V (Room Controller) | pcs | 11 |  | 0 | #1-#8 rooms · #9-#10 B3 · #11 B2 (corridor ceiling)
7 | EVV049 | SmartHub connectia (Raspberry Pi) | pcs | 1 |  | 0 | in the rack
8 | EOP006 | LVNS server | pcs | 1 |  | 0 | in the rack
9 | EOP002 | Nurse iPad 10.9" (Wi-Fi) | pcs | 1 |  | 0 | nurse desk · receives and resolves all calls
10 | EVV038 | ionTab licences + MDM | lot | 1 |  | 0
[5 empty] | Subtotal | 0
NETWORK AND RACK
11 | GAB | 42U communications rack (with 230 V) | pcs | 1 |  | 0 | end of corridor · Wi-Fi network by others
[5 empty] | Subtotal | 0
CABLING (model lengths + reserve, see Cables sheet)
12 | BUS | 2x2x0.8 mm bus (RS485 + 12-48 VDC) – PCB daisy chain (green) | m | 45 |  | 0 | model length + reserve
13 | UTP | Radial call cable device → PCB (green) | m | 190 |  | 0 | model length + reserve
14 | NET | Cat6 – SmartHub and LVNS server in the rack (red) | m | 5 |  | 0 | model length + reserve
15 | PWR | Power 3x2.5 mm² LSZH – rack only (blue) | m | 10 |  | 0 | model length + reserve
[5 empty] | Subtotal | 0
INSTALLATION
16 | INS-EQ | Mounting and connection of devices (panels, lights, pull-cords, PCB/LCom/BoxCom, tablets) | pcs | 52 |  | 0 | sum of all devices
17 | INS-CAB | Cable pulling in tray / ceiling, conduit and labelling | m | 250 |  | 0 | sum of all cabling
18 | INS-RACK | Rack assembly: switch, server, patch panel, UPS | lot | 1 |  | 0
19 | INS-PRG | Programming, commissioning and testing of every call point | lot | 1 |  | 0
20 | INS-ASB | As-built drawings and handover protocol | lot | 1 |  | 0
[5 empty] | Subtotal | 0
TRAINING
21 | TR-USR | Training for nursing / reception staff (daily use) | session | 2 |  | 0 | per shift
22 | TR-TEC | Technical training for maintenance / IT | session | 1 |  | 0
23 | TR-DOC | User and technical manuals (ES / EN / HE) | lot | 1 |  | 0
[5 empty] | Subtotal | 0
WARRANTY AND SERVICE
24 | GAR-EQ | Equipment warranty | year | 1 |  | 0 | as per manufacturer
25 | GAR-INS | Installation and workmanship warranty | year | 1 |  | 0
26 | MNT | Preventive maintenance and technical support (optional) | year | 1 |  | 0 | optional
[5 empty] | Subtotal | 0
HOSPITAL INFRASTRUCTURE (other trades)
27 | CT | Bed-head unit (O₂/air/vacuum, 230 V, RJ45, light) | pcs | 12 |  | 0 | 1 per bed
28 | — | Hospital bed + bedside cabinet + overbed table | pcs | 12 |  | 0
29 | — | U-shaped curtain track + fire-retardant curtains | lot | 4 |  | 0 | 1 per room
30 | — | Fixed / folding grab bars, shower seat | lot | 1 |  | 0 | B2 · B3
[5 empty] | Subtotal | 0
[5 empty] | Subtotal | 0
[5 empty] | Contingency | 0
[5 empty] | Subtotal incl. contingency | 0
[5 empty] | VAT (Israel) | 0
[5 empty] | TOTAL INCL. VAT | 0
Note on C6: Assumption: 5 % contingency. Adjust.
Note on C7: Israel VAT 18 % (since 01/2025). Set 0 % if not applicable.

## Cables
Cables
Reserve | 15%
Round up to | 5
Code | Description | Calculated (m) | Qty to buy (m)
BUS | 2x2x0.8 mm bus (RS485 + 12-48 VDC) – PCB daisy chain (green) | 35.8 | 45
UTP | Radial call cable device → PCB (green) | 161 | 190
NET | Cat6 – SmartHub and LVNS server in the rack (red) | 1.8 | 5
PWR | Power 3x2.5 mm² LSZH – rack only (blue) | 5.5 | 10

## Notes
Notes
1. Quantities taken from the project's 3D/2D model.
2. Cable lengths measured along the drawn routes; reserve and rounding are set in the Cables sheet.
3. Unit prices: fill in from the current Advantecnia / supplier price list (yellow cells).
4. 5 % contingency and 18 % VAT are editable assumptions (BOQ sheet, cells C6 and C7).
5. Preliminary quantities: verify on site before purchasing.