## Proposal
IMPORTANT – THIS IS NOT A REAL, FORMAL OR OFFICIAL PROPOSAL. It is an example only, prepared to support the offer that Orad must make to the end client.
Technical-Commercial Proposal – Example Proposal 1
Example Proposal 1 · Yoseftal Medical Center, Eilat · ORAD-YOS-EX1-2026-10 · 26/10/2026
1 ward section · 8 rooms × 3 beds · 8 en-suite bathrooms · 1 nurse station
|  | Orad discount (all items) |  |  |  | 30%
Currency: EUR · prices ExW (Incoterms 2020) · VAT not included
| Code | Description | Location | Qty | List price | Disc. | Net price | Amount
Nurse call devices – rooms and bathrooms |  |  |  |  | Category discount | 30%
| ACE82012153 | Bed call panel, red, with bayonet socket for pear handset 1-module PVC panel 114×114 mm, red 30 mm domed button with Braille and reassurance LED, female bayonet pigtail for the pear handset. One per bed, at the bed head (h 1.15 m). | 3 per room × 8 rooms | 24 | 95.9 | 30% | 67.13 | 1611.12
| ACE39060521 | Pear handset, 1 capacitive button, 2 m cable, bayonet Hand-held call button for the patient in bed, antibacterial plastic, reassurance LED and Braille; plugs into the bed call panel. | 1 per bed × 24 beds | 24 | 80.07 | 30% | 56.05 | 1345.2
| ACE82082137 | Presence / cancel panel, green 1-module PVC panel, green 30 mm button: the nurse marks presence and cancels the call. Inside each room by the door and outside each bathroom (h 1.35 m). | 8 rooms (by the door) + 8 bathrooms (outside) | 16 | 95.9 | 30% | 67.13 | 1074.08
| ACE40320303 | SOS mechanism with pull-cord (toilet / shower) 45×45 SOS mechanism on 80×80 European frame, 2 m nylon cord with ergonomic handle and anti-break device. One beside the toilet and one beside the shower (h 1.85 m). | 2 per bathroom (toilet + shower) × 8 | 16 | 63 | 30% | 44.1 | 705.6
| EOP004 | Corridor light, 3 colours (red · green · blue) LED door light with LEGRAND frame, cover and diffuser, above each room and bathroom door (h 2.42 m): red = call, green = presence, blue = code blue. | 8 rooms + 8 bathrooms (above the door) | 16 | 183.67 | 30% | 128.57 | 2057.12
|  | Subtotal – Nurse call devices – rooms and bathrooms | [5 empty] | 6793.12
Digital Room – Shabbat system (controllers, server, nurse station, licences) |  |  |  |  | Category discount | 30%
| ECV169 | Room controller PCB 24 pins, 12–48 V DC Up to 5 alarm inputs and 1 light output, in a junction box near the call point; daisy-chained on a 2×2×0.8 mm bus (RS485 + 12–48 V) from the SmartHub. 2 per 3-bed room, 1 per bathroom. | 2 per room + 1 per bathroom = 24 | 24 | 163.27 | 30% | 114.29 | 2742.96
| EVV049 | SmartHub connectia – head of the controller bus 19" 1U, 230 VAC input, 12–48 V DC output to the PCB bus; gateway to the IP network. One per ward section. | 1 per ward section | 1 | 1326.53 | 30% | 928.57 | 928.57
| EOP006 | LVNS server – local server & notifications Central server: manages every call, presence and alarm, sends them to the nurse iPads and keeps the history. One server serves all ward sections. | 1 server for the whole system | 1 | 2653.06 | 30% | 1857.14 | 1857.14
| EOP002 | Nurse station iPad 10.9" Wi-Fi 128 GB Nurse station outside the rooms: receives and resolves every call (Consola Monitor), sets or cancels presence in any room. Includes case, desk stand and charger. | 1 per nurse station | 1 | 936.51 | 30% | 655.56 | 655.56
| EVV038 | ionTab client/server licences + MDM (1st year included) Licences for the nurse iPad, the room PCBs, the SmartHub and the LVNS server, plus device management (MDM). Per ward section – quantity per PCB to be confirmed by ACE BG. | 1 per ward section | 1 | 406.13 | 30% | 284.29 | 284.29
|  | Subtotal – Digital Room – Shabbat system (controllers, server, nurse station, licences) | [5 empty] | 6468.52
Network & rack (switch, rack, UPS, Wi-Fi) |  |  |  |  | Category discount | 30%
| ACESWP24P | PoE switch 24 ports, managed Network switch in the rack: connects the SmartHub, the LVNS server and the Wi-Fi access point. | 1 per ward section rack | 1 | 410.66 | 30% | 287.46 | 287.46
| RACK9U | Wall rack 19" 9U with PDU, shelf and cable management Local supply (Orad) – estimate, to confirm. | 1 per ward section | 1 | 457.14 | 30% | 320 | 320
| UPS1500 | UPS 1500 VA online, rack mount Local supply (Orad) – estimate, to confirm. Keeps the system running during power cuts. | 1 per rack | 1 | 685.71 | 30% | 480 | 480
| WAP6 | Wi-Fi 6 access point, PoE, at the nurse station Local supply (Orad) – estimate, to confirm (if the hospital Wi-Fi is not used). | 1 per nurse station | 1 | 314.29 | 30% | 220 | 220
|  | Subtotal – Network & rack (switch, rack, UPS, Wi-Fi) | [5 empty] | 1307.46
|  | TOTAL SYSTEM (net to Orad) | [5 empty] | 14569.1
Not in the total – to confirm / optional
| SW-UPD | Software licences & updates – years 2 and 3 Year 1 is included in the system price (EVV038). Years 2–3: price to be confirmed by ACE BG. | 2 years | 2 |  |  |  | 0
| SW-MAINT | Optional: full remote support Optional service, 15% of the system total per year. | per year | 1 |  | 15% |  | 2185.365
| SPARES | Recommended spare parts See the separate spare-parts quotation. | 1 set | 1 |  |  |  | 1066.64
Summary
|  | Nurse call devices – rooms and bathrooms | [5 empty] | 6793.12
|  | Digital Room – Shabbat system (controllers, server, nurse station, licences) | [5 empty] | 6468.52
|  | Network & rack (switch, rack, UPS, Wi-Fi) | [5 empty] | 1307.46
|  | TOTAL SYSTEM (net to Orad) | [5 empty] | 14569.1
Not included
• Cabling, conduits, trays, junction boxes and back boxes (quantities in the BOQ)
• Installation, programming and on-site commissioning (local subcontractor / Orad)
• Hospital LAN / Wi-Fi integration works
• VAT, transport, customs and insurance
• Code Blue (can be added on request)
Comments
• Prices: ACE BG net prices to Orad, from the reference quotation for Yoseftal (Shabbat option). List price = net ÷ 0.70 (Orad discount 30%).
• Yellow = local supply estimate (rack, UPS, Wi-Fi AP) or price to be confirmed.
• Maximum number of PCBs per SmartHub bus to be confirmed by ACE BG for this scope.
• Validity: 30 days. Warranty: 1 year on manufactured products. Delivery time: to be confirmed with the order.
• Discount note: ACE BG prices to Orad already include a 30% discount on the list price (List price = net ÷ 0.70). Version 1.30.
Yellow = editable input (discount, estimate or price to confirm). Blue = input, black = formula. Change the yellow discount cell to apply a new discount to everything, or change a category / line discount.
EXAMPLE ONLY – NOT A REAL, FORMAL OR OFFICIAL PROPOSAL